CLOVER 10

Handheld Video Magnifier

User Guide

Rev 2.0

Contents

Precaution	1

Quick Start	3

Overview	3

Compositions	5

Accessories	8

Battery Charging	9

Stand	11

External Assist Stand (Optional)	13

Operation	14

Power Button	14

Beep Type	15

Beep Control	15

Camera Switch	15

Magnification	16

Color Mode	17

Brightness	19

Lighting	20

Focus Control	21

Image Freeze	22

Image Storage	23

Image Playback	24

Line & Mask	27

Find	27

Live Panning	29

Power Frequency	30

Memory	31

HDMI Output	32

Auto Power Saving	33

Restore	33

Hardware Reset	34

Troubleshooting	35

The unit does not turn on	35

Black screen or White screen	35

The image does not change	36

Smudges or blemishes on the screen	36

Specifications	37

# Precaution

Before using Clover 10 for the first time, you must charge its battery. Refer to Battery Charging section for more details.

Always operate the unit in a place where the temperature is between 10°C and 35°C (50°and 95°F).

Keep the unit away from wet areas, rain, and other sources of liquids. Do not submerge the unit.

Do not attempt to repair the unit. Any unauthorized attempt will void the product warranty. Please contact your distributor.

Use the supplied power adapter and battery.

# Quick Start

## Overview

Thank you for purchasing Clover 10.

Clover 10 is a portable 10-inch video magnifier with excellent Full HD image quality, which assures the real-time crisp image while moving or changing magnification continuously from 2X to 32X. The integrated dual camera system provides three viewing ways to fulfill the daily use. For example, Near View to read books or newspapers; Far View to see posters or blackboards; Mirror View to dress up.

Besides the common zooming and color mode function, there are also various practical functions such as Reading Line & Mask, Image Freeze and Storage, Find and Navigation.

## Compositions

| 1. LCD Screen | 2. Front Camera | 3. Enhanced Color |
| --- | --- | --- |
| 4. Zoom Out | 5. Joystick | 6. Zoom In |
| 7. True Color |  |  |

| 8. Rear Camera | 9. LED Lights | 10. Stand |
| --- | --- | --- |

| 11. Power Button | 12. Power Port |
| --- | --- |
| 13. HDMI Port | 14. Locate Button |
| 15. Freeze Button | 16. Camera Switch |
| 17. FW Updating Port | 18. Focus Control |

## Accessories

| User Manual | Power Adapter | HDMI Cable |
| --- | --- | --- |
| Cleaning Cloth | Carrying Case |  |

## Battery Charging

1. Connect the supplied power cable to the power adapter.

2. Connect the supplied power cable to the unit’s power port.

3. Plug the power adapter into an AC outlet and start charging.

Power LED

- Charging when the unit is off, the power LED will be green, and then auto turn off when fully charged. A full Charge will take about 5.5 hours and will support the unit working for about 4.5 - 5.5 hours.

- Charging when the unit is on, the power LED will be green.

- When the unit is on without charging, the power LED will be green when the battery level is high and will be red when the battery is dying. As long as the red light starts to flash, it is a signal to charge the battery.

- You can also check the battery status by pressing the power button. Refer to the following instructions for more details.

## Stand

With integrated fold-away writing stand at an ergonomic viewing angle, clover 10 can satisfy users’ desire to read more and for a long time in a comfortable position. Also, there’s enough space under the camera to work the pen. Unfold the stand as the following figure shows.

The rear camera is designed to auto rotate according to the status of the stand, so that both the Near View and Far View are supported on this unit.

## External Assist stand (Optional)

For people who has handwriting request, there is a new upgraded external assist stand option, which can recognize the device and enter into the handwriting mode with all lights on when the device is inserted. More information please refer to the stand manual.

# Operation

## Power Button

- Press and hold the power buttonfor 3 seconds to turn on/off the unit.

- When the unit is on, press the power buttonto check the current status of battery capacity, freeze mode, viewing mode, and beep.

- Press other buttons when holding the power button will make combo functions. Refer to the following instructions for more details.

## Beep Type

- Unit on/off beep

- Button beep

- Threshold beep

- Low battery beep

## Beep Control

- Press and hold the Power buttonand press the True Color buttonto turn on/off the beep.

## Camera Switch

- Turn the Camera Switchleft to open the rear camera for Near View and Far View.

- Turn the Camera Switch right to open the front camera for Mirror View.

## Magnification

- Press the Zoom In buttonto increase magnification, and hold the Zoom In buttonfor continuous increase.

- Press the Zoom Out buttonto decrease magnification, and hold the Zoom Out buttonfor continuous decrease.

- Magnification range when standing on table:

Back/Front Camera: 2x – 32x

## Color Mode

A color mode is a predefined color combination. Full-color mode is suitable for viewing pictures or maps.

Enhanced Color mode is combined with different foreground and background colors which can help to read more clearly.

- Press the Enhanced Color buttonto go through the Enhanced-color combination list.

- Press the True Color buttonto quickly return to the Full-color mode.

- Press and hold the Enhanced Color buttonfor 3 seconds to enter the Customized Color Page, use the Joystickand the True color buttonto activate/inactivate the enhanced color combination.

- No operations for 8 seconds or short press the Power button button or long press the Enhanced Color button to exit the Customized Color Page.

- Press the Enhanced Color button and Zoom In/Out can adjust contrast color. Factory default value will be highlighted green.

## Brightness

Brightness of the screen can be adjusted in both Full-color mode and Enhanced-contrast mode.

- Press and hold the Power buttonand press the Zoom In/Out / to increase/decrease the brightness. Factory default value will be highlighted green.

## Lighting

There are 4 lighting statuses can be applied on the device.

| Lights all off | Lights all on |
| --- | --- |
| Left light on, Right light off | Left light off, Right light on |

- Press and hold the Power buttonand click the Freeze buttonto turn on/off the two LED lights simultaneously (The operation is compatible for all firmware version)

- Press and hold the Power buttonand press the Focus Control buttonto switch from 3 lighting statuses: Left light on, Right light on and All lights on/off (based on the latest setting).

## Focus Control

- Click the Focus Control button to lock the focus, which is especially useful when handwriting.

- Click the button again to unlock the focus.

- Long press the  can trigger the focus manually.

## Image Freeze

- Freezing an image on the LCD screen can be useful for further reading. At this status, magnification as well as color change is still allowed for a better experience.

- Press the Freeze buttonto lock the active image.

- Press the Freeze buttonagain to unlock.

- For the locked image with increased magnification, you can use the Joystickto navigate around freely.

## Image Storage

- Long press the Freeze buttonfor 3 seconds to store the current image. You can save up at least 1000 images, 2GB.

- Note: When the saved images reach to the maximum value (Memory Full Icon appears), you will need to delete some pictures first in Playback mode before the next successful saving.

## Image Playback

- Long press the True button  for 3 seconds to enter the image gallery.

- Slide the Joystick/Zoom In/Out / to select the target image from the Four Grids Thumbnail Image Gallery. Press the Freeze buttonor True button  to view the target image. You can view and navigate the image with the same operations described in the Image Freeze section. Press the Freeze buttonto return to the Gallery.

- When playback the image, press the Power buttonto pop up the Previous and Next Page icon. Press the Zoom In/Out /to turn pages without exiting to the Galley.

- When in the Gallery, press the Locate buttonto select the target Image which needs to be deleted. Press the Locate buttonagain to confirm, or any other buttons to cancel.

- When in the Gallery, hold the Locate buttonfor 5 seconds to delete all the images. Press the Locate buttonagain to confirm, or any other buttons to cancel.

- Long press the Freeze buttonfor 3 seconds to activate the multi-selection. Use the Joystickor Zoom In/Out / and True Color buttonto choose the target pictures. And you can delete the selected pictures at once.

- Press the Locate button  will pop up confirm box. Press the True Color button to confirm deletion and the Enhanced Color button to cancel.

- Long press the Freeze buttonagain or long press the True Color button or short press the Power buttonto cancel the multi-selection.

- Long press the True Color button for 3 seconds or short press the Power button to exit the Playback mode.

- In the gallery, keeping sliding down/up the Joystickfor 3 seconds to start turning the pages down/up quickly.

## Line & Mask

Press the Locate buttonto switch among different Reading Line & Mask modes. The Live-panning function is still available.

Hold the Locate buttonand then click the Enhanced color button to change the line color.

- One horizontal line

- Press and hold the Locate buttonand slide the Joystick  up/down to change the position of horizontal reading line; left/right to decrease/increase the width of the line.

- One vertical line

- Press and hold the Locate buttonand slide the Joystick  left/right to change the position of vertical reading line; up/down to increase/decrease the width of the line.

- Horizontal mask

- Press and hold the Locate buttonand slide the Joystick  up/down to change the position of the masks; left/right to narrow/enlarge the distance between the masks.

- Vertical mask

- Press and hold the Locate buttonand slide the Joystick  left/right to change the position of the masks; up/down to enlarge/narrow the distance between the masks.

Note: Press the Power buttoncan close the line/mask quickly.

## Find

- Long press the Locate button, a red location box would appear.

- Move and change the content in the location box, aiming it at the part you want to view by sliding the Joystick, and then press the Locate button again, the image will return back to the previous magnification with the refreshed content.

- Press the Zoom Inand Zoom Out at the same time to help re-center the location box.

Live Panning

- In the state of amplification, users can slide the Joystick  when in the real-time video mode to navigate the image for seeing more without moving the unit.

- Press the Zoom Inand Zoom Outbuttons at the same time to help you re-center the image (Red Dot appears).

- Two ways to quickly jump to the start/end point: 1) press and hold the Power button, and slide the Joystick. 2) quickly push the Joystickin the same direction twice.

- The navigation speed can be adjustable in three levels: slow, medium (the factory default) and fast. Push the Joystickto one side and press the Zoom In/Out /to increase/decrease the reading speed.

## Power Frequency

Power Frequency provides the options of 50HZ and 60HZ power frequency according to customer region. Long press the Power button + Enhanced color buttonfor 3 seconds can adjust the frequency.

## Memory

The unit will memorize automatically the most recent settings such as color mode, magnification, beep when power is off normally and apply them when the unit is on next time.

## HDMI Output

Connect the unit to your external display device with HDMI cable for a better and larger viewing (resolution: 1080P).

## Auto Power Saving

If no operation and no move for over 3 minutes, the unit will turn off automatically.

- Press and hold the Power buttonand click the Enhanced Color button to turn on/off the Auto Power Saving function manually (factory default is on).

## Restore

Long press the Freeze buttonand the Locate buttonat the same time for 5 seconds to reset to the default factory settings. Auto power off when reset is done.

## Hardware Reset

When the unit is hung or can’t be turned off normally, press the Power+ Freeze button+ True Color buttonat the same time to trigger the hardware reset.

# Troubleshooting

## The unit does not turn on

- Connect to the power adapter

- Charge the battery

- Hardware Reset

## Black screen or White screen

- If the unit is sitting on a flat pure surface, the screen may be black or white. Lift the unit off to ensure change happens

- Adjust the magnification

- Adjust the color mode

- Restart the unit

- Hardware Reset

## The image does not change

- If the unit enter the freeze mode or playback mode, the image will not change

- Press power button to check the current status

- Restart the unit

- Hardware Reset

## Smudges or blemishes on the screen

- Adjust brightness

- Clean the LCD screen and camera

Please contact your local distributor if problem is not solved.

# Specifications

| Display: | 10.1 inch LCD screen, HD |
| --- | --- |
| Camera: | Auto Focus, 13 MP |
| Size: | <= 245*200*22mm |
| Weight: | <= 700 grams |
| Color Mode: | True color,customized color mode; 16 Integrated enhanced color modes |
| Magnification: | 2x – 32x |
| Battery: | Integrated rechargeable Li battery; 4.5-5.5 hours continuous use after 5.5-hour charge |
| Power Adapter: | Input: 110-240V; Output: 5V/3A |